## Slide 1
232Th and 238U Concentrations measured by the ICP-MS in PKU
Ying YUAN Hao QIAO Siguang WANG
School of Physics, Peking University

## Slide 2: Stainless Steel
Samples | Th-232 （pg/g)* | U-238 （pg/g)*
P4TC | 378 17 | 312  24
P4TF | 606 129 | 1087 100
P4TN | 30350 | 75571
P4TM | 30976 | 60086
P4TO | 5610 | 1026100
P4TP | 87258 | 76487
*Note：averagestdev
Surface treatment
Dissolved in aqua regia(王水）
Diluted
Analyze(Using standard addition method or dilution)
39K of P4TO: 27、26 and 28(ng/g) for 3 samples

## Slide 3: MicroMegas (GEM)
Detection Limits：
0.24 pg 232Th/g (0.97 μBq 232Th/kg)
0.14 pg 238U/g (1.73 μBq 238U/kg)
Surface treatment
Microwave Digestion with concentrated HNO3(spiked with 230Th and 235U)
Loaded on the Extraction Chromatography Column
Analyze
 | Preliminary | 
Sample | 232Th pg/g | 238U pg/g
mm 1 | 36 16 | 2737
mm 2 | 34 13 | 30842
mm 3 | 11825 | 29234
Ca. 0.088 g

## Slide 4: Micromegas “BIPO”
Sample | Weight g | 232Th pg/g | 238U pg/g
1 | 0.09796 | 46 | 287
2 | 0.10556 | 43 | 266
Over simplified method summary:
Sample into microwave digestion instrument
Dissolve in nitric acids
Dilute and analyze

## Slide 5: Copper(Single crystal copper)
flow rate ：0.8 mL/min
Ca. 0.85 g
Surface treatment
Dissolved in 8 M HNO3 (spiked with 230Th and 235U)
Loaded on the Extraction Chromatography Column
Analyze
Detection Limits：
0.035 pg 232Th/g (0.14 μBq 232Th/kg)
0.073 pg 238U/g (0.90 μBq 238U/kg)
 | Preliminary | 
Sample | 232Th pg/g | 238U pg/g
Cu 1 | 0.220.07 | 0.370.08
Cu 2 | 0.28 0.14 | 0.370.07
Cu 3 | 0.270.07 | 0.590.19
Cu 4 | 0.160.07 | 0.5760.033
Cu 5 | 0.140.07 | 0.540.15

## Slide 6: 3D printer raw material “Feinpolyamide PA 2200 for EOSINT P”
Sample | Weight g | 232Th pg/g | 238U pg/g
1 | 0.09522 | 283 | 3586
2 | 0.09369 | 290 | 3386
Over simplified method summary:
Sample into microwave digestion instrument
Dissolve in nitric acids
Dilute and analyze